# WIKI System Start UP

Table of Contents

Procedure information	2

1.	Reference documents	3

2.	Bill of Materials needed	3

3.	Terminologies	3

4.	Zeroing Process Manometer	4

a.	Zeroing using manometer Tuning Screw driver	4

b.	Zeroing using Vat valve (indirect approach)	5

c.	Confirming zeroing	6

5.	External data logger	7

a.	Settings for external triggering.	7

6.	Appendix	10

a.	Installing VAT VALVE Software (CPA)	10

# Procedure information

| Procedure # | VS_WIKI_1007_REVA |
| --- | --- |
| Initial Issue Date | 06-12-2021 |
| Initial Edited by | Behnam Behziz |
| Last Edit Date | 06-12-2021 |
| Last Edited by | Behnam Behziz |
| Affected systems | All |
| Expected Execution Time | 1 hours |
| Number of Operators Needed | 1 Person |
| Operator requirements | Technician level 1 |
| Electrical Hazard | Non. |
| Mechanical Hazard | Non. |
| Chemical Hazard | Non. |
| Safety Control | Non. |

# Reference documents

- NA

# Bill of Materials needed

- NA

# Terminologies

- Manometer: gas pressure sensor

# Zeroing Process Manometer

- Zeroing the manometer is possible through manometer tuning screw or through Vat valve program.

- For each system once both “a,b” needs to be done. After that, the zeroing can be done by using “a” only.

## Zeroing using manometer Tuning Screw driver

    - Ensure the manometer is being turned on for more than 60 minutes. This ensures the PCBA is in thermal equal prior to hand-tuning the manometer.

    - Use HMI. Open up the Vat valve fully. Stop any gas flow to chamber.

    - Ensure the full-scale manometer is reading lower than 1 uTorr.

    - Measure the voltage between pin 2 and 12 (2 positive and 12 negative) using FLUKE voltmeter;

    - Record the voltage and pressure from HMI in table below.

    - Hand-tune the tuning screw in front of the manometer to reach to 0 as close as possible. (Absolute error smaller than 0.05V)

    - Measure new voltage between pin 2 and 12(2 positive and 12 negative); Record it in table below.

    - Record the pressure from HMI in table below.

    - Disconnect the FLUKE setup and connect back the cable.

|  | Before adjustment | After Adjustment |
| --- | --- | --- |
| Voltage (pin 2&12) |  |  |
| Pressure |  |  |

## Zeroing using Vat valve (indirect approach)

- Ensure the manometer is being turned on for more than 60 minutes. This ensures the PCBA is in thermal equal prior to hand-tuning the manometer.

- Use HMI. Open up the Vat valve fully. Stop any gas flow to chamber.

- Locate VAT cable from red bin in electrical tool box.

- Use the cable, and connect USB side to your computer and connect the Dsub side to service port of the VAT valve.

- Open CPA program to communicate to the VAT VALVE. If CPA program is not installed, please follow appendix (‎5).

- The program will try to scan the VAT valve. Connect to correct port.

- From top left, click on the key symbol and ensure it is set to “Remote”.

- Navigate to Sensor > setup and enable zeroing.

- Navigate to Sensor > Zero and click on Zeroing.

- Note that the absolute value of offset value is less than 0.05V.

- From top left, click on the key symbol and ensure it is set to “Local”.

- Disconnect service cable.

## Confirming zeroing

    - Ensure the manometer is being turned on for more than 60 minutes. This ensures the PCBA is in thermal equal prior to hand-tuning the manometer.

    - Use table below as reference and fill out the readings. Consult with Engineer to ensure the numbers are acceptable.

    - For instance, for case 3, open the vat valve, flow gas to 10 SCCM, wait until the reading of flow stabilizes. Record the fine manometer reading. Record the full-scale manometer reading. Record the voltage reading from pin 2 and 12.

    - For instance, for case 7, flow gas to 10 SCCM, wait until reading of flow stabilizes. Set the pressure setting of vat valve to 3 mTorr. Wait until the reading of pressure stabilizes. Record the fine manometer reading. Record the full-scale manometer reading. Record the voltage reading from pin 2 and 12.

| Case | Flow (SSCM) | VAT VALVE Pressure setpoint (mTorr) | Fine Manometer (mTorr) | Full Scale Manometer (Torr) | Voltage Reading (V) |
| --- | --- | --- | --- | --- | --- |
| 1 | 0 | Open |  |  |  |
| 2 | 5 | Open |  |  |  |
| 3 | 10 | Open |  |  |  |
| 4 | 15 | Open |  |  |  |
| 5 | 20 | Open |  |  |  |
| 6 | 6 | 1.5 |  |  |  |
| 7 | 10 | 3 |  |  |  |
| 8 | 15 | 6 |  |  |  |

# External data logger

## Settings for external triggering.

- Ensure the datalogger is stopped by pressing stop if it is not grayed out.

- Select Capture settings to adjust the settings for external signal

- Select “Trigger Alarm Settings” Tab

- Set Trigger start condition to “External”

- Set Trigger stop condition to “External”

- Set the repeat to “ON”

- Apply

- Ok. This forces the datalogger listen to external signal generated from HMI to triggering start/stop of datalogging.

- Press start.

- Ensure the system goes to “Armed” mode. Now the datalogger is waiting for external signal to start recording or end recording.  When HMI starts a recipe, it will send an external trigger to start recording and when the recipe completed, it will end the recording and saves it for future review.

# Appendix

## Installing VAT VALVE Software (CPA)

  - Please install VAT Valve CPA software from this link.

  - Follow the installation instruction from the link above.